## Sheet1
No. | Prompt | Expected Response | MCP
1 | Why cant i communicate with my Trace 1300? | Based on the TRACE 1300 documentation, loss of communication between the GC and the PC is usually due to basic power or LAN configuration problems. The manuals recommend checking the following: Confirm the instrument is powered correctly Make sure the TRACE 1300 power cable is firmly connected at both the GC and the mains outlet. Verify the wall outlet itself is working (for example, by testing with another device). If the power system appears faulty, you are instructed to contact your local Thermo Fisher Scientific customer support organization. Check the LAN/Network cabling Ensure the LAN cable is properly connected to the TRACE 1300 LAN (RJ45) connector and to the PC or network switch. Use the correct cable type: Red reversed patch cable for a direct PC-to-TRACE 1300 connection (used for initial setup or when not on the site LAN). Grey standard patch cable for connection to the site LAN. Verify TCP/IP settings and LAN configuration Confirm that the IP address, netmask, and port number configured on the PC exactly match the settings of the TRACE 1300. The GC is shipped with a factory IP address that may not match your lab network, so your LAN administrator must provide: An appropriate IP address (e.g., 192.168.xxx.xxx) A netmask (e.g., 255.255.255.0) A communication port (e.g., 2551) If the LAN settings on the GC are unknown or incorrect, you can use the reset button on the LAN module (press for at least 5 seconds with a small tool) to restore default LAN parameters before reconfiguring. Restart and verify configuration/handshaking Restart the TRACE 1300. Verify that GC methods and configuration are correct in the control software. Confirm that the GC handshaking parameters and instrument configuration in the software match the actual hardware. Environmental and site-LAN considerations The preinstallation guide notes that Thermo Fisher only checks communication on a mini-hub or LAN switch; responsibility for connecting to and configuring your site LAN lies with your local IT/network team. If your GC communicates correctly on a small local switch but not on the building LAN, you likely need help from your LAN administrator. If, after checking power, cables, network settings, and configuration/handshaking, the TRACE 1300 still cannot communicate with the PC, the manuals direct you to contact Thermo Fisher Scientific technical support for further diagnosis. | If you can’t communicate with your TRACE 1300, the most common causes are related to power, network (LAN), or configuration issues: LAN connection problems: The LAN cable may not be properly connected, the wrong cable is being used, or there is no LAN activity at the GC. IP/TCP-IP mismatch: The computer’s network settings do not match the GC’s IP address, netmask, or port configuration. Instrument not restarted: Communication issues can persist until the GC is restarted after configuration changes. Power or startup state: If the GC is not fully powered on, not ready, or has a faulty power supply, communication with the software will fail. Software/configuration issues: Methods may not download or the software may not recognize the GC if the instrument configuration or handshaking parameters are incorrect. Typical corrective actions include verifying power and LAN cabling, confirming the TCP/IP settings match between the PC and GC, restarting the instrument, and checking that the GC is correctly configured in the software. For more information, please read: • mt_poc_test_31715002_-Rev_F-_TRACE_1300_and_TRACE_1310_Gas_Chromatographs_Hardware_Manual%20%282%29.pdf • mt_poc_test_31715001 - Rev F - TRACE 1300 and TRACE 1310 Gas Chromatographs Preinstallation Requirements Guide.pdf
2 | What is the newest version of firmware available for the Trace 1300? | The latest documented firmware for the TRACE 1300 Series GC is: TR1300_07.05.22_2022_09_09 with Main firmware version: Main_07.05.22.hex This release is announced in Product Support Bulletin 2022_0041, which describes it as a “new GC Trace 1300 Series Firmware (TR1300_07.05.22_2022_09_09)” and lists it as the latest set of module firmwares (Main, Oven, SSL, PTV, detectors, AUX_Temp, etc.). An earlier 2022 bulletin (HERMES_Main_07.05.19) and older packs (48, 39, 37) are superseded by this 07.05.22 package. | The newest firmware available for the TRACE 1300 is GC Trace 1300 Series Firmware version 07.05.22, released as part of Product Support Bulletin 2022_0041 on November 2, 2022. This release includes improvements to the GC booting phase when a hydrogen sensor is installed and fixes issues such as negative PTV ramps. Updating to this firmware is strongly recommended, especially for systems equipped with a hydrogen sensor. For more information, please read: • mt_poc_test_PSB%20GCMS_2022_0041%20New%20GC%20Trace%201300%20Series%20Firmware%20%28TR1300_07.05.22_2022_09_09%29.pdf
3 | Where can i find maintenance videos for the Trace 1600 series GC? | For the TRACE 1600 Series GC, preventive maintenance (PM) training videos are available as internal e-learning content. Thermo Fisher employees can access them via Microsoft Stream at the channel: “GC TRACE 1600 / 1300 series PM ¦ Microsoft Stream” This e-learning includes step-by-step PM videos for SSL, PTV, FID, TCD, ECD, NPD, FPD, and gas leak detection. In addition, the TRACE 1610/1600 touch screen itself has a Video page in the Maintenance menu that shows step-by-step videos for instrument tasks and maintenance procedures, depending on your GC configuration. | Maintenance videos for the TRACE 1600 series GC are available through Thermo Fisher’s internal preventive maintenance e-learning resources. These include step-by-step videos covering common PM procedures such as SSL, PTV, FID, TCD, ECD, NPD, FPD, and gas leak detection. Thermo Fisher employees can access these videos via Microsoft Stream at the internal link titled “GC TRACE 1600 / 1300 series PM”. In addition, TRACE 1600/1610 instruments with a touch screen provide built-in maintenance videos directly on the HMI, accessible from the Maintenance or Video pages, depending on the instrument configuration. For more information, please read: • mt_poc_test_PM%20Trace%20GC%201600%20Series%20and%20Trace%20GC1300%20Series.html • mt_poc_test_MI-317000-0034%20-%20TRACE%201600%20User%20Guide.pdf
4 | How do i turn off the beeping noise from the Trace 1600 series GC? | The beeping noise on the TRACE 1600 Series GC is triggered only during an alarm condition. When an alarm occurs, both the Power and Ready LEDs blink and a beep sounds. The GC also automatically cuts power to heaters and gases. The system does not provide a user-configurable setting to disable the beeper. To stop the beeping, you must clear the alarm condition by resetting the instrument: How to stop the beeping: Identify and resolve the cause of the alarm (e.g., temperature over-limit, gas supply, detector fault, etc.). Power-cycle the GC: Turn the TRACE 1600 off, Wait a few seconds, Then turn it on again. Once the alarm state is cleared, the beeping stops automatically. | On the TRACE 1600 series GC, the beeping noise is an alarm buzzer that sounds when the instrument detects a malfunction or abnormal operating condition. The beep cannot be manually disabled. It will stop only when the underlying alarm condition is resolved. What you can do: Identify and resolve the alarm condition (for example, temperature over-limit, gas issue, detector not ready, or other hardware faults). Once the issue is fixed, reset the alarm: For TRACE 1600: power the GC off, then power it back on. For TRACE 1610: after correcting the issue, go to the Diagnostics menu on the touch screen and press Reset to reboot the system. After the alarm is cleared and the GC returns to a normal state, the beeping will stop automatically. The buzzer is a safety feature and is intentionally tied to alarm states rather than a user-configurable setting. For more information, please read: • mt_poc_test_MI-317000-0034%20-%20TRACE%201600%20User%20Guide.pdf • mt_poc_test_MI-317000-0033%20-%20TRACE%201600%20Hardware%20Manual.pdf
5 | What is the part number for the iconnect fitting for the trace 1600 series ssl? | The iConnect fitting (iConnect Column Lock) used with the TRACE 1600 Series SSL injector has the part number: MI-190000-0028 – iConnect Column Lock | The iConnect fitting (Column Lock) used for the TRACE 1600 series SSL (Split/Splitless) injector has the following part number: iConnect Column Lock — Part Number: MI-190000-0028 This iConnect fitting is used for column installation and retention in the SSL and SSLBKF injectors on TRACE 1600/1610 systems. For more information, please read: • mt_poc_test_MI-317000-0035%20%28Rev.%20C%29%20TRACE%201600%20Series%20Spare%20Parts%20Guide.pdf • mt_poc_test_MI-317000-0033%20-%20TRACE%201600%20Hardware%20Manual.pdf
6 | How often does the Trace 1300 series ecd need to be wipe tested? | The TRACE 1300 Series ECD must be wipe tested every six (6) months. This is a regulatory requirement for radioactive Ni-63 detectors and applies to all TRACE 1300 ECD units. Additional notes from the documents: A wipe test is also required before shipping an ECD. For users in the United States, the 6-month interval is mandated; other countries must follow their respective regulatory agencies. | This six-month interval is a regulatory requirement (notably in the United States) to check for radioactive contamination from the Ni-63 source. In addition, the ECD must be wipe tested before shipment, and if a detector fails a wipe test, the licensee must take the appropriate regulatory action. For regions outside the U.S., equivalent local regulatory agency requirements apply, but the 6-month interval is the standard reference. For more information, please read: • mt_poc_test_PSB_GCMS_2019_0001%20-%20TRACE%20ECD%20Return%20Procedure%20and%20Radioactive%20Material%20Warning.pdf • mt_poc_test_31715002_-Rev_F-_TRACE_1300_and_TRACE_1310_Gas_Chromatographs_Hardware_Manual%20%282%29.pdf
7 | Can the TCD filament be replaced? | No — on the TRACE 13x0 / 16x0 TCD, the filament itself is not replaceable as an individual spare part. When the filaments are old, contaminated, or damaged, the entire TCD module must be replaced. | No. The TCD filament cannot be replaced individually. If the filament is old, contaminated, or broken, the entire TCD detector module must be replaced. Thermo Fisher documentation explicitly states that the filaments are not serviceable as standalone parts. Filament-related issues (low response, high noise, high standing current, or detector not going ready) all require replacement of the complete TCD module rather than a filament swap. For more information, please read: • mt_poc_test_TCD%20Module%20Troubleshooting%20guide%20.pdf • mt_poc_test_MI-317000-0034%20-%20TRACE%201600%20User%20Guide.pdf
8 | Can the Trace 1300 series be upgraded to a trace 1600 series? | Yes. A TRACE 1300 Series GC can be upgraded to a TRACE 1610 GC, which is part of the TRACE 1600 Series family. The TRACE 1600 Series GC Service Manual includes a specific procedure titled “Upgrading TRACE 1300 Series GC to a TRACE 1610 GC”, which describes step-by-step how to convert a TRACE 1300 into a TRACE 1610 by replacing the front panel and associated electronics. This is treated as a formal hardware upgrade, not just a firmware change, so it should be done using the proper upgrade kit and procedure via Thermo Fisher service. In practice, you should: Use the official upgrade procedure and hardware described in the service documentation. Work with your Thermo Fisher Scientific service representative to identify the correct upgrade kit and arrange the upgrade. | Yes — a TRACE 1300 series GC can be upgraded to a TRACE 1610 (TRACE 1600 series), but this is a hardware upgrade, not a software or firmware-only conversion. Thermo Fisher provides a defined upgrade procedure that involves replacing key hardware components, most notably: Installing the TRACE 1610 front panel (HMI) Replacing the electronic module and associated cabling Performing the required mechanical and electrical installation steps as documented in the service manual This upgrade effectively converts the instrument into a TRACE 1610, adding the touch-screen interface and TRACE 1600–series functionality. It is a service-level operation and not intended as a simple field retrofit without proper parts and procedures. For more information, please read: • mt_poc_test_MI-317000-0063%20TRACE%201600%20Series%20Service%20Manual.pdf • mt_poc_test_MI-317000-0014%20%28Rev.%20B%29%20-%20AI%20AS%201610%20Spare%20Parts%20Guide.pdf
9 | where does the auxillary power module plug in on the Trace 1300? | On the TRACE 1300, the Auxiliary modules (Auxiliary Gas module / Aux Temperature-Cryo module) connect in two places: GC BUS connection to the TRACE 1300: Use the supplied 15-pin cable. Plug the connector labeled “GC BUS” on the auxiliary module into a BUS interface on the back of the TRACE 1300 GC (power module). The module only works if this GC BUS connector is attached to the BUS connector on the TRACE 1300. AC power connection: The auxiliary module has its own AC Input connector on the front of the module. Plug the module’s power cable into this AC Input and then into the wall outlet (it does not draw power directly from the GC). | On the TRACE 1300, the Auxiliary Power Module (used for options such as the Aux Temperature/Cryo Module or NPD Thermionic Source Power Module) is not powered through the GC bus. Instead: It plugs directly into AC mains power using its own power cable, connected to the AC Input connector on the front of the auxiliary module, and then to the wall outlet. The module must be powered at the same line voltage as the TRACE 1300 GC. Control and communication with the GC occur via the GC Bus cable, which connects from the auxiliary module to a GC Bus interface on the back of the TRACE 1300. Important distinctions: Aux Temperature/Cryo Module and NPD Power Module → powered by their own AC input Auxiliary Gas Module, Generic Detector Interface, and Analog Output Interface → powered directly from the GC via the GC Bus For more information, please read: • mt_poc_test_31715002_-Rev_F-_TRACE_1300_and_TRACE_1310_Gas_Chromatographs_Hardware_Manual%20%282%29.pdf • mt_poc_test_TRACE%201300_1310%20Service.pdf
10 | What are the differences in the Trace 1300 series and the Trace 1600 series? | Feature / Aspect TRACE 1300 Series TRACE 1600 Series (e.g. TRACE 1610) User Interface / Controls Basic — older panel/controls. Upgraded — newer front-panel, potentially touchscreen (depending on config), more modern electronics. Upgrade Path Original/base model. Designed to be an upgrade — a TRACE 1300 can be converted to TRACE 1610 by replacing the front panel and associated electronics via an official upgrade kit. Spare Parts & Module Structure Uses older module layout (SSL, PTV, detectors) per 1300 spec sheet. Uses a revised hardware architecture as per 1600 Series — potentially easier maintenance, newer components and spare-parts set. Maintenance / Serviceability Standard for its time. Updated — as the 1600 Series Service Manual is newer, it may include improved maintenance procedures, support for newer options/accessories, and updated spare-parts handling. Firmware & Software Support Firmware updates are available (latest: TR1300_07.05.22). As part of the upgrade path, the 1600 Series will leverage newer firmware/OS/hardware — likely benefitting from longer-term support. Accessory/Module Compatibility Uses “legacy” interfaces (e.g. GC BUS, AC input for aux modules). Some accessories and modules (detectors, aux, etc.) may differ — though upgrade kits may allow reuse if compatible. | The TRACE 1300 series and TRACE 1600 series gas chromatographs share a common modular philosophy, but the TRACE 1600 series introduces significant enhancements in usability, safety, and operational efficiency. Key differences: User Interface TRACE 1300: Single-button operation with full control through the CDS (designed for minimal local interaction). TRACE 1600: Enhanced local control options. TRACE 1600: single-button local interface. TRACE 1610: advanced high-resolution touch screen with direct instrument control, diagnostics, maintenance videos, and alerts. Column Handling TRACE 1300: Traditional ferrule-based column installation. TRACE 1600: Proprietary iConnect™ Column Lock, enabling tool-free, faster, and more reliable column installation with reduced leak risk. Maintenance & Usability TRACE 1300: Modular injectors and detectors with offline maintenance capability. TRACE 1600: Builds on this with: On-screen maintenance video guides Consumable tracking and alerts for GLP compliance Faster troubleshooting and reduced downtime Safety & Gas Management TRACE 1300: Requires optional hydrogen sensor kit for hydrogen carrier gas. TRACE 1600: Supports advanced HeSaver-H2Safer™ technology, improving helium conservation and hydrogen safety integration. Productivity & Design TRACE 1600 is more compact and optimized for workflow efficiency, with improved uptime and operational profitability while retaining compatibility with existing injectors, detectors, and autosamplers. In summary, the TRACE 1300 is optimized for routine, CDS-driven operation with minimal local interaction, while the TRACE 1600 is a next-generation platform focused on enhanced user experience, safety, and reduced total cost of ownership. For more information, please read: • mt_poc_test_PS52260_E%20-%20TRACE%201300%20Series%20Gas%20Chromatograph%20Specification.pdf • mt_poc_test_Product%20Specification%20-%20TRACE%201600%20Series%20Gas%20Chromatograph.PDF